## Slide 1: Introduction fMRI based on FUNCTIONAL Magnetic Resonance Imaging by Huettel et al.

## Slide 2: MRI
1971 Raymond Damadian proved that magnetic resonance could be used to help detect diseases by the different nuclear magnetic relaxation times between tissues and tumors.
1973 Paul Lauterbur demonstrated magnetic resonance imaging on test tubes.
1977 Raymond Damadian completed the first whole body MRI scanner which he called "Indomitable."
1977 Raymond Damadien performed the first whole body scan using Indomitable. It lasted four hours and 45 minutes to complete.
1992 functional MRI was developed (Ogawa and Kwong)

## Slide 3: Overview
fMRI Scanner
- static magnetic field, gradient coils, safety etc.
MR Signal generation
- magnetic moment, relaxation
MR Image formation
- slice selection, frequency and phase encoding
MR Contrast
T1, T2 and T2*
BOLD-signal
- neurovascular coupling, characterisation of the signal

## Slide 4: fMRI scanner
Siemens, 3T Trio
Static magnetic field (0.5-14 T)
Radiofrequency coils
Gradient coils
Shimming coils
Computer hardware and software

## Slide 5: Static magnetic field
Generated by Helmholtz Coils
- Currents are parallel (same direction)
- Field along MRI axis
Superconducting electromagnets
-(0.5-11 Tesla, -261°C)
field uniformity
- Homogentiy in space and time

## Slide 6: Radiofrequency coils
Defined by function:
- Transmit/receive coil
- Transmit only coil
- Receive only coil
Defined by geometry
- Surface coil (High sensitivity but limited coverage)
- Volume coil (low sensitivity but uniform coverage)
- Phased-array coil (High sensitivity, near-uniform coverage)

## Slide 7: Gradient and shimming coils
A number of different electromagnetic coils are used to generate the
B0-field and the gradients in x-, y-, and z-direction. Shimming coils
generate compensatory magnetic fields correcting for inhomogenitie
in the B0 field.
solenoid
solenoid

## Slide 8: Gradient coils
A Maxwell pair, two loops with opposing currents, which generate magnetic field gradients along Z direction, i.e., the main magnetic field.
Golay pair which allows generation of magnetic field gradients in the X and Y direction, i.e., perpendicular to the main magnetic field.

## Slide 9: Computer hardware and software

## Slide 10: MR Safety
The magnet is never switched off!
Subjects are excluded having metal implant, pacemaker, aneurysm clip, tattoos,
foreign metal objects, neurostimulator etc. or being pregnant.

## Slide 11: MR Signal generation

## Slide 12: Nuclear Spin
A spinning proton and a spinning bar magnet have angular momentums (J) and magnetic moments (μ). The angular momentums are generated by the spinning masses; the magnetic moment for the spinning proton is due to the electric current, which is induced by the rotating charge. The magnetic moment for the bar magnet comes from the movement of its internal magnetic field

## Slide 13: Magnetic moment
In a magnetic field:
protons can take either
high- or low-energy states
In the absence of a strong
magnetic field:
the spins are oriented randomly

## Slide 14: Precession
The movement of a rotating proton or magnet within a magnetic field is similar to the movement of a top in Earth’s gravitational field, i.e., . the axis of spin itself (spin axis) wobbles around the main axis of the magnetic or gravitational field (precession axis).

## Slide 15: Net Magnetization
Small B0 produces small net magnetization
Thermal motions try to randomize alignment of proton magnets
Larger B0 produces larger net magnetization, lined up with B0
At room temperature, the population ratio is roughly 100,000 to
100,006 per Tesla of B0

## Slide 16: Excitation of a Spin System
More spins in an external magnetic field will be at a low-energy level
Excitation pulse is applied and some spines will jump to the high-energy state
High-energy spins will return to the low energy state (relaxation)

## Slide 17: Excitation of a Spin System
The delivery of an excitation pulse (B1) causes the longitudinal magnetization (M) to be tipped into the transverse plane. The longitudinal net magnetization would trace out a complex wobbling path as it rotates downward to the transverse plane

## Slide 18: Signal Reception
Before Excitation
After Excitation
Excitation tips the net magnetization down into the transverse plane,
where it can generate current in detector coils (i.e., via induction).

## Slide 19: Relaxation
In absence of external B1, the net magnetization will go back being
aligned with static field B0 — this is called relaxation
Part of net magnetization perpendicular to B0 shrinks; this part of
is called transverse magnetization (T2) and provides the detectable
RF signal
Part of net magnetization parallel to B0 grows back; it is called
longitudinal magnetization (T1) which is not directly detectable

## Slide 20: T2-Relaxation
T2: Intrinsic decay of the transverse magnetization
- Usually 50-100 ms in the brain [shortens with bigger B0]
T2*: Overall decay of the observable RF signal
- Usually about half of T2 in the brain [i.e., faster relaxation]

## Slide 21: T1-Relaxation
T1: Relaxation of net magnetization back to alignment with B0
- Usually 500-1000 ms in the brain [lengthens with bigger B0]

## Slide 22: MR Image formation

## Slide 23: Magnetic gradients
Gradient coils generate varying magnetic field so that spins at different
location precess at frequencies unique to their location, allowing us to
reconstruct 2D or 3D images.
X gradient Y gradient Z gradient
x
y
z
x
z
z
x
y
y

## Slide 24: Slice selection
Excite and receive net magnetization in a thin (2D) slice of the subject
- RF signal must come from this slice
- Reduce dimension from 3D down to 2D

## Slide 25: Pulse sequence for slice selection
To select a slice, an excitation pulse is delivered through the radiofrequency coil (RF; a head coil in fMRI) with a magnetic gradient. Each line of a pulse sequence diagram indicates a separate component of the scanner hardware (x-axis indicating time, y-axis indicating the strength of that
Component).

## Slide 26: Frequency encoding
Deliberately make magnetic field strength B depend on location within slice
- Frequency of RF signal will depend on where it comes from
- Breaking signal into frequency components will provide localization information

## Slide 27: Frequency and phase encoding
During readout (image acquisition) period, turning on gradient field is called
frequency encoding --- using a deliberately applied nonuniform field to make
the precession frequency depend on location.
Before readout (image acquisition) period, turning on gradient field is called
phase encoding --- during the readout (image acquisition) period, the effect
of gradient field is no longer time-varying, rather it is a fixed phase accumulation
determined by the amplitude and duration of the phase encoding gradient.

## Slide 28: Pulse sequence with frequency encoding
To create a one-dimensional map within a slice, a second magnetic gradient must be applied during data acquisition. Note that data acquisition usually occurs with some delay after excitation for T1- or T2-sensitive images.

## Slide 29: Pulse sequence with frequency and phase encoding
To create a two-dimensional image of a given slice, two independent magnetic gradients must be used within that slice, i.e., a phase encoding gradient is applied before the frequency-encoding gradient, so that there is time for phase differences among spins to accumulate along that direction.

## Slide 30: k-space

## Slide 31: MR Contrast
Contrast = difference in RF signals (emitted by water protons)
between different tissues

## Slide 32: Proton-density
Technique: use very long time between RF (large TR) and very short
delay between excitation and readout window (short TE)
- useful for anatomical reference scans
~ 1 mm resolution

## Slide 33: T1
Technique: use intermediate timing between RF (intermediate TR) and
very short TE
- useful for anatomical reference scans
~1 mm resolution

## Slide 34: T2
Technique: use large TR and intermediate TE
- useful for anatomical and functional studies
- 1mm resolution for anatomical scans or 3 mm functional scans

## Slide 35: EPI and T2*
Methods shown earlier take multiple RF shots to readout enough
data to reconstruct a single image
Each RF shot gets data with one value of phase encoding
If gradient system (power supplies and gradient coil) are good
enough, can read out all data required for one image after one
RF shot (signal is available in 2*T2* [80 ms])

## Slide 36: BOLD-signal affects T2
Increasing Blood Oxygenation
Decreasing Relaxation Time
T2
T1
Thulborn et al., 1982

## Slide 37: BOLD-signal
Oxygenated blood in tube leads to little signal change, either on
spin- or gradient-echo images
Deoxygenated blood leads to large susceptibility effects on gradient-
echo images
Ogawa 1990

## Slide 38: BOLD Endogenous Contrast
Blood Oxyenation Level Dependent Contrast
  Deoxyhemoglobin is paramagnetic, oxyhemoglobin is less so.
  Magnetic susceptibility of blood decreases linearly with increasing oxygenation
Oxygen is extracted during passage through capillary bed
  Arteries are fully oxygenated
  Venous (and capillary) blood has increased proportion of deoxyhemoglobin
  Difference between oxy and deoxy states is greater for veins  BOLD sensitive to venous changes

## Slide 39: BOLD-signal
Kwong 1992
 VISUAL 
 MOTOR 

## Slide 40: BOLD-signal
Ogawa 1992
High-field (4T) in humans
Patterned visual stimulation at 10 Hz
Gradient-echo (GRE) pulse sequence used
  Surface coil recorded

## Slide 41: Neurovascular coupling
Neurovascular coupling refers to the relationship between local neural activity
and subsequent changes in cerebral blood flow (CBF)

## Slide 42: Mechanism of BOLD Functional MRI
Brain activity
Oxygen consumption
Cerebral blood flow
Oxyhemoglobin
Deoxyhemoglobin
Magnetic susceptibility
T2*
MRI signal intesity

## Slide 43: BOLD-signal and neuronal activity
BOLD response reflects pooled
local field potential activity
(Logothetis et al, 2001)

## Slide 44: Time resolution BOLD-signal
Delay in the HDR: Hemodynamic activity lags neuronal
activity by 4-6 sec

## Slide 45: Non-linearity BOLD-signal

## Slide 46: Variability of the BOLD-signal
Across Subjects
Across Sessions
Across Brain Regions
Across Stimuli

## Slide 47
Introduction SPM

## Slide 48
SPM
Open Source academic freeware (under GPL)
Documented and informally supported
Requirements:
  MATLAB: 7.1 (R14SP3) to 7.9 (R2009b) no toolboxes (EEG/MEG requires ‘signal’)
  Supported platforms (MEX files):
  File Formats:
    Images: NIfTI-1 (& Analyze, DICOM)
    Surface meshes: GIfTI
    M/EEG: most manufacturers (with FieldTrip’s fileio)
Linux (32 and 64 bit)
Windows (32 and 64 bit)
Mac Intel (32 and 64 bit)

## Slide 49: SPM Interface

## Slide 50
Terminology of fMRI
Structural (T1) images:
- high resolution
- to distinguish different types of tissue
Functional (T2*) images:
- lower spatial resolution
to relate changes in BOLD signal
to an experimental manipulation
 Time series:
A large number of images that
are acquired in temporal
order at a specific rate
t

## Slide 51
Overview
Normalisation
Statistical Parametric Map
Image time-series
Parameter estimates
General Linear Model
Realignment
Smoothing
Design matrix
Anatomical reference
Spatial filter
Statistical Inference
RFT
p <0.05

## Slide 52
Pre-Processing
Normalisation
Statistical Parametric Map
Image time-series
Parameter estimates
General Linear Model
Realignment
Smoothing
Design matrix
Anatomical reference
Spatial filter
Statistical Inference
RFT
p <0.05

## Slide 53
Realignment
Even small head movements can be a major problem:
  Movement artefacts add up to the residual variance and reduce sensitivity.
  Data may get completely lost if sudden movements occur during a single volume
  Movements may be correlated with the task performed
  -> Minimising movements is one of the most important factors for ensuring
  good data quality

## Slide 54
Realignment
Each volume in the time-series is realigned
  Using rigid-body registration
    Assumes that brain shape doesn’t change
  Least squares cost function
Realigned images must be resliced for analysis
  Not necessary if they will be normalised anyway

## Slide 55
Realignment
A 3D rigid body transform is defined by:
  3 translations - in X, Y & Z directions
  3 rotations - about X, Y & Z axes
The order of the operations matters
Translations
Pitch
about x axis
Roll
about y axis
Yaw
about z axis

## Slide 56
Realignment
Minimising mean-squared difference works for intra-modal
registration (realignment)
Simple relationship between intensities in one image, versus
those in the other
  Assumes normally distributed differences

## Slide 57
Realignment

## Slide 58
Slice-time-correction
Re-sampling can introduce interpolation errors especially
tri-linear interpolation
Slices are not acquired simultaneously rapid movements not
accounted for by rigid body model
Image artefacts may not move according to a rigid body model
  image distortion
  image dropout
  Nyquist ghost
Functions of the estimated motion parameters can be modelled
as confounds in subsequent analyses

## Slide 59
Normalization
Inter-subject averaging
  Increase sensitivity with more subjects
  Fixed-effects analysis
  Extrapolate findings to the population as a whole
  Mixed-effects analysis
Make results from different studies comparable by aligning them to
standard space e.g. using the MNI template

## Slide 60
Normalization
305
T1
PD
T2
SS
“Canonical” images in MNI space
The Talairach atlas is not
exactly accurate for inter-
preting coordinates from
SPM analyses

## Slide 61
Normalization
Non-linear registration
Minimise mean squared difference from template image(s)
Affine registration

## Slide 62
Normalization
The first part is a 12 parameter affine transform
  3 translations
  3 rotations
  3 zooms
  3 shears
Fits overall shape and size
Algorithm simultaneously minimises
  Mean-squared difference between template and source image
  Squared distance between parameters and their expected
  values (regularisation)

## Slide 63
Normalization
Algorithm simultaneously minimises
  Mean squared difference between
  template and source image
  Squared distance between parameters
  and their known expectation
Deformations consist of a linear combination of
smooth basis functions
These are the lowest frequencies of a 3D
discrete cosine transform

## Slide 64
Smoothing
Why would we deliberately blur the data?
  - Suppress noise and effects due to differences in
  anatomy by averaging over neighbouring voxels
  - Achieve better spatial overlap, enhanced sensitivity,
  and greater validity of statistical assumptions
  - Reduce the effective number of multiple
  comparisons when correcting results using Random
  Field Theory
How is it implemented?
Convolution with a 3D Gaussian kernel, of specified
Full-width at half-maximum (FWHM) in mm

## Slide 65
Before convolution
Convolved with a circle
Convolved with a Gaussian
Smoothing is done by convolution.
Each voxel after smoothing effectively becomes the result of
applying a weighted region of interest (ROI).
Smoothing

## Slide 66
Many interesting and clinically important questions might relate
to the shape or local size of regions of the brain
For example, whether (and where) local patterns of brain
morphometry help to:
- Distinguish schizophrenics from healthy controls
- Understand plasticity, e.g. when learning new skills
- Explain the changes seen in development and aging
Voxel-Based Morphometry

## Slide 67
Alzheimer’s Disease example
Baseline Image
Standard clinical MRI
1.5T T1 SPGR
1x1x1.5mm voxels
Repeat image
12 month follow-up rigidly registered
Subtraction image

## Slide 68
Alzheimer’s Disease example
Some changes are apparent in this patient...
Some might be noise or misregistration etc.
Perhaps confounding biological effects like hydration changes etc.
But some might genuinely reflect underlying AD pathology...
Would changes generalise from the patient to the disease?
It is appealing to try a “second-level” SPM analysis of structural data
variation over subjects
E.g. AD vs. healthy, male vs. female or a correlation with IQ

## Slide 69
Pre-processing for VBM

## Slide 70
Tissue segmentation
High-resolution MRI reveals fine structural detail in the brain, but not all of
it reliable or interesting
  Noise, intensity-inhomogeneity, vasculature, …
MR Intensity is usually not quantitatively meaningful
  fMRI time-series allow signal changes to be analysed statistically,
  compared to baseline or global values
Regional volumes of the three main tissue types: gray matter, white
matter and CSF, are well-defined and potentially very interesting
  Other aspects (and other sequences) can also be of interest

## Slide 71
Tissue segmentation
Structural image, GM and WM segments, and brain-mask (sum of GM and WM)

## Slide 72
Tissue segmentation
Intensities are modelled by a Gaussian Mixture Model (AKA Mixture Of Gaussians)
With a specified number of components
Parameterised by means, variances and mixing proportions (prior probabilities for components)

## Slide 73
Tissue Probability
Tissue probability maps (TPMs) are used instead of the proportion of voxels in each Gaussian as the prior.
ICBM Tissue Probabilistic Atlases (kindly provided by the International Consortium
for Brain Mapping, John C. Mazziotta and Arthur W. Toga.

## Slide 74
Aligning the tissue probability maps
Initially affine-registered using a multi-dimensional form of mutual information
Iteratively warped to improve the fit of the unified segmentation model to the data
  Familiar DCT-basis function concept, as used in normalisation

## Slide 75
Segmentation

## Slide 76
DARTEL
Rigid average
Template 6
Template 1

## Slide 77
Modulation
Multiplication of the warped (normalised) tissue intensities so that their regional or global volume is preserved
  Can detect differences in completely registered areas
Otherwise, we preserve concentrations, and are detecting mesoscopic effects that remain after approximate registration has removed the macroscopic effects
  Flexible (not necessarily “perfect”) registration may not leave any such differences
1 1
2/3 1/3 1/3 2/3
1 1 1 1
Native
Modulated
Unmodulated

## Slide 78
Total Intracranial Volume (TIV/ICV)
“Global” integrated tissue volume may be correlated with
interesting regional effects
  Correcting for globals in this case may overly reduce sensitivity
  to local differences
  Total intracranial volume integrates GM, WM and CSF, or
  attempts to measure the skull-volume directly
    Not sensitive to global reduction of GM+WM (cancelled out by
    CSF expansion – skull is fixed!)
  Correcting for TIV in VBM statistics may give more powerful
  and/or more interpretable results

## Slide 79
Statistical Parametric Mapping…
group 1
group 2
voxel by voxel modelling
–
parameter estimate
=
gCBF
r
CBF
x
o
o
o
o
o
o
x
x
x
x
x
g..
a
k
1
a
k
2
z
k
1

## Slide 80
Segment
Normalise
VBM in pictures

## Slide 81
Modulate
VBM in pictures

## Slide 82
VBM in pictures
Voxel-wise statistics